## Slide 2: 1000, 100, 10, 1 ppb
BTEX in Air
with
Calidus GS-T
“Limit of quantification”

## Slide 3: Objectives
Separate and identify ppb-level BTEX components in air
  benzene
  toluene
  ethylbenzene
  meta- & para-xylene (not separated)
  ortho-xylene
Establish repeatability at varying concentrations with an RSD of <5% per component
    Concentration levels : 1000ppb, 100ppb, 10ppb, 1ppb
Establish linearity for each component across the four orders of magnitude
Explore possibility of parts per trillion level measurements

## Slide 4: Calidus GS-T Configuration

## Slide 5: Calidus GS-T Configuration
The GS-T sample pump consistently loads the trap module for an adjustable length of time
Sample is drawn from the top of the instrument for ambient air measurements
  To meet the objectives of this test, sample was drawn from Tedlar® bags via a needle at the end of a length of PEEKTM tubing

## Slide 6: Calidus Method
The vacuum pump was set to concentrate sample in the trap module for 100 seconds
  Approximately 50 mL of air is drawn with a sample pump time of 100 seconds
The injection was splitless
Sample was concentrated and held isothermally on the trap at 35°C during the 100s sample collection
After this time, the trap module was quickly ramped at 20°C/s to 250°C
The concentrated contents of the 250°C trap module were “injected” onto the Q-Bond column module where the BTEX components were separated and measured

## Slide 7: BTEX Sample Preparation
BTEX was injected into several zero air-filled Tedlar bags in order to achieve the desired concentrations of 1ppb, 10ppb, 100ppb, and 1000ppb
All concentrations refer to benzene in BTEX
  All other compounds are present at about 25% lower concentrations
  For example, at the 100ppb level, individual component concentrations were:
    benzene: 100ppb
    toluene: 85ppb
    ethylbenzene: 74ppb
    o,p,m-xylene: 74ppb each

## Slide 8
OBJECTIVES 1 and 2 – Identification and Repeatability

## Slide 9: BTEX – 1000ppb
Unknown hydrocarbons suspected to be inherent in the Tedlar bags at concentrated levels is noticeable in three large non-interfering peaks

## Slide 10: 8 Runs Overlaid - 1000ppb

## Slide 11: 8 Runs Area Repeatability – 1000ppb
Run Number | benzene | toluene | ethylbenzene | p+m-xylene | o-xylene
1 | 283225 | 273159 | 231392 | 564746 | 282758
2 | 283194 | 275630 | 231541 | 567166 | 281034
3 | 279843 | 272464 | 228479 | 559647 | 277142
4 | 285006 | 277541 | 232522 | 571944 | 285854
5 | 285418 | 278216 | 234280 | 571646 | 286893
6 | 286189 | 279068 | 233980 | 574886 | 287709
7 | 280800 | 271309 | 228156 | 558204 | 276504
8 | 284837 | 278308 | 233433 | 573476 | 286632
Average | 283564 | 275712 | 231723 | 567714 | 283066
Std. Dev. | 2260 | 3026 | 2348 | 6354 | 4457
%RSD | 0.80 | 1.10 | 1.01 | 1.12 | 1.57

## Slide 12: BTEX – 100ppb
Dip in baseline due to large plug of air in the sample hitting the FID
Only two unknown peaks from this Tedlar bag

## Slide 13: 8 Runs Overlaid - 100ppb

## Slide 14: 8 Runs Area Repeatability – 100ppb
Run Number | benzene | toluene | ethylbenzene | p+m-xylene | o-xylene
1 | 30700 | 27253 | 24660 | 54206 | 29361
2 | 30885 | 27351 | 24760 | 54577 | 29536
3 | 30749 | 27230 | 24674 | 54258 | 29289
4 | 30709 | 27140 | 24626 | 54108 | 29317
5 | 30504 | 27174 | 24399 | 53940 | 29072
6 | 30694 | 27376 | 24562 | 54278 | 29344
7 | 30542 | 27138 | 24528 | 53760 | 29105
8 | 30301 | 26683 | 24189 | 53154 | 28769
Average | 30635 | 27168 | 24550 | 54035 | 29224
Std. Dev. | 180 | 215 | 182 | 431 | 235
%RSD | 0.59 | 0.79 | 0.74 | 0.80 | 0.81

## Slide 15: BTEX – 10ppb
Dip in baseline due to end of splitless time at 100s

## Slide 16: 8 Runs Overlaid - 10ppb

## Slide 17: 8 Runs Area Repeatability – 10ppb
Run Number | benzene | toluene | ethylbenzene | p+m-xylene | o-xylene
1 | 2796 | 2453 | 2091 | 4413 | 1866
2 | 2687 | 2513 | 2097 | 4501 | 1863
3 | 2644 | 2462 | 2117 | 4444 | 1910
4 | 2629 | 2422 | 2091 | 4432 | 1890
5 | 2713 | 2468 | 2126 | 4482 | 1915
6 | 2707 | 2490 | 2153 | 4480 | 1921
7 | 2667 | 2491 | 2159 | 4528 | 1938
8 | 2618 | 2463 | 2087 | 4416 | 1871
Average | 2683 | 2470 | 2115 | 4462 | 1897
Std. Dev. | 57 | 28 | 29 | 42 | 28
%RSD | 2.14 | 1.13 | 1.36 | 0.94 | 1.49

## Slide 18: BTEX – 1ppb

## Slide 19: 8 Runs Overlaid - 1ppb
Baseline rise and drift due to column bleed as column approaches max temp

## Slide 20: 8 Runs Area Repeatability – 1ppb
Run Number | benzene | toluene | ethylbenzene | p+m-xylene | o-xylene
1 | 230 | 264 | 143 | 511 | 118
2 | 223 | 267 | 149 | 511 | 117
3 | 210 | 261 | 142 | 510 | 123
4 | 215 | 248 | 150 | 479 | 122
5 | 219 | 253 | 146 | 504 | 112
6 | 218 | 264 | 140 | 507 | 113
7 | 227 | 266 | 157 | 527 | 118
8 | 219 | 266 | 139 | 500 | 114
Average | 220 | 261 | 146 | 506 | 117
Std. Dev. | 6 | 7 | 6 | 13 | 4
%RSD | 2.87 | 2.64 | 4.13 | 2.67 | 3.34

## Slide 21
OBJECTIVE 3 – Linearity

## Slide 22: Linearity (1, 10, 100, 1000ppb)
Chart: Benzene Linearity
| 1 | 10 | 100 | 1000
Benzene | 226.6074 | 2682.6269 | 30675.4133 | 283564.1875
Chart: Ethylbenzene Linearity
| 1 | 10 | 100 | 1000
EthylBenzene | 151.0467 | 2228.9956 | 24579.4389 | 231722.7625
Chart: p+m-Xylene Linearity
| 1 | 10 | 100 | 1000
p+m Xylene | 522.1717 | 4886.2171 | 54093.9456 | 567714.3875
Chart: o-Xylene Linearity
| 1 | 10 | 100 | 1000
o xylene | 131.707 | 3031.6568 | 29250.0322 | 283065.65
Chart: Toluene Linearity
| 1 | 10 | 100 | 1000
Toluene | 283.2281 | 2470.3219 | 27200.9844 | 275711.7625

## Slide 23
OBJECTIVE 4 – Parts Per Trillion Measurements

## Slide 24: BTEX – 500ppt

## Slide 25: BTEX – 250ppt

## Slide 26: Overlay - 1ppb, 500ppt, 250ppt (zoomed to see benzene and toluene)
1ppb
500ppt
250ppt

## Slide 27: Objectives Met
All components exhibited an RSD of 5% or better (much better in most cases) at all four concentration levels of 1, 10, 100, and 1000ppb
All components exhibited linear responses through four orders of magnitude with R squared values of 0.9999 or better (1.0 in most cases)
Parts per trillion levels of BTEX components were still detectable and integrateable by the Chromperfect software

## Slide 28: Further Studies
Some of the later eluting peaks would likely benefit from a lower final temperature hold on the Q-Bond
  The method in this study brought the column to its maximum temperature limit of 300°C in order to increase speed of analysis
  This high temperature hold caused the baseline to rise towards the end of the run which made the integration of some peaks (especially at the lower ppb and ppt levels) very difficult
Seeing even lower ppt levels than this experiment showed would be possible by pulling air through the sample pump for longer than 100s