Location of the potentiometer.

Location of test point on

PCB P/N B74674

1- Introduce zero gas to the analyzer and let it run for 15 minutes.

2- Then connect the voltmeter negative to TP1 and positive to TP2

3- Adjust the pot so that the DMM read <10mv.\

4- Zero calibrate

5- Span calibrate.